# BIOCAGE Cervical Cage: Computer Simulation Study

What we are doing, what we have learned, and what it means for the design

Status as of 2026-10-25  |  Human C3-C4 finite-element study  |  Author: Yi Liu (austonliuyi)  |  English version (Chinese version follows)

How to read the source marks: a blue superscript [n] is a published reference (list in Section 8). (Decision: ...) is a choice agreed with the principal investigator (PI) and recorded with its date in the project's decision log; "Q47" etc. are the question numbers in that log. (Comm.: ...) is an instruction from, or a conversation with, the PI. (Protocol: Pn) points to test procedure Pn in Section 3.3. "auto (no reply)" means the PI did not answer within about 30 minutes, so the analyst decided using the PI's standing priority rule (1 closest to real human physiology, 2 supported by literature or standards, 3 the analyst's reasoned recommendation, 4 saves time without losing much accuracy) and told the PI afterwards. Technical terms are explained in the glossary (Section 2.4).

# 1. Summary in one page

The BIOCAGE is a small 3D-printed polymer cage placed between two neck vertebrae after the damaged disc is removed. Its window holds a lab-grown replacement disc (the "biodisc"). The cage only has to hold the space open and stay in place while the biodisc matures inside the body (Comm.: PI 2026-10-23). Earlier bench tests with real cages showed that the main problem is the cage moving out of place (dislocation) (Comm.: PI 2026-10-21). We therefore use a validated computer model of a human neck segment to test many cage variants and find which features keep the cage stable and strong enough.

The computer model is trustworthy for bending: it reproduces the published reference model and experimental data within about 10 % [1,2,3,4,5](Decision: Validation passed, PI 2026-10-20).

A big error was found and fixed: the model first needed 50-100 times more force to spread the vertebrae than real cadaver spines do. The cause was a ring of leftover disc tissue; it is now treated as removed, as in surgery [21,22](Decision: PI 2026-10-22).

Full-size, shape-matched cages (design A) are far better than narrow flat cages with a soft pad (design B): 108 of 120 design-A variants carry at least 150 N; none of 24 design-B variants does.

The soft surface layer is the weak spot in every test, and friction alone may not hold the cage: under everyday bending with the lowest plausible friction, every design that got that far is predicted to slip.

The test must match the device's job: the PI replaced the original motion test (made for permanent artificial discs) with an everyday-load test (Decision: PI 2026-10-23 12:00); a numerical problem in backward bending still has to be solved before it can be used for all designs.

# 2. Background for new readers

## 2.1 The clinical problem

The neck (cervical spine) has seven vertebrae, named C1 to C7 from the head down. Between neighbouring vertebrae sits an intervertebral disc: a tough fibrous outer ring (annulus) around a gel-like centre (nucleus). The disc is attached to the flat top and bottom surfaces of the vertebrae (endplates). Behind the disc are two small sliding joints (facet joints), and ligaments run along the front (ALL) and back (PLL) of the spine. This study looks at the level between the 3rd and 4th vertebrae (C3-C4).

When a disc degenerates and presses on nerves, surgeons remove it through the front of the neck. They then either fuse the two vertebrae with a rigid cage, or insert an artificial disc that keeps the joint moving (ACDR, anterior cervical disc replacement) (Decision: Surgery 2026-10-20).

## 2.2 The BIOCAGE idea

A cage 3D-printed from PLCL, a slowly degrading medical polymer. Mixing a stiff grade (PLCL 85:15) and a rubber-like grade (PLCL 50:50) gives regions that are stiffer or softer [9,10,11,28].

A central window (60 % of the cage area) holds the biodisc, a tissue-engineered disc [12](Decision: Cage geometry 2026-10-20).

Purpose: temporary support and fixation while the biodisc matures in the body. It is not meant to be a permanent moving disc or a fusion device (Comm.: PI 2026-10-23).

Main risks: the cage slides or tips out of place (dislocation); the cage breaks; the cage sinks into the bone (subsidence); the space collapses before the biodisc is ready.

## 2.3 Why computer simulation, and how it works

Finite-element (FE) simulation divides the bones, tissues and cage into many small pieces (elements). It then solves the physics equations to predict forces, deformation, sliding and damage. This lets us test hundreds of cage variants under identical conditions, which bench tests cannot do quickly. The model must first be checked against real measurements (validation) (Protocol: P1). Each simulated case follows the same sequence: build the spine segment, perform the "surgery", insert the cage, then apply a test load (outline in Section 3.3).

We use FEBio, a free research FE program for biomechanics [25], version 4.13. We fixed five small software faults in its source code and checked that the fixes do not change correct results (Decision: Patch verification 2026-10-20). We also added a graphics-card (GPU) solver that makes each run about 2-4 times faster [26].

## 2.4 Glossary

| Term | Meaning in this report |
| --- | --- |
| C3-C4 segment / FSU | The 3rd and 4th neck vertebrae with the disc, joints and ligaments between them (functional spinal unit). |
| Endplate | The top or bottom bony surface of a vertebra that the cage rests on. |
| Annulus remnant | The ring of outer disc tissue that may be left after the surgeon removes the disc. |
| ACDR | Anterior cervical disc replacement: removing the disc from the front and inserting a device that keeps motion. |
| Biodisc | The lab-grown (tissue-engineered) disc held in the cage window; very soft (0.03 MPa). |
| PLCL hard / soft phase | Stiff PLCL 85:15 (250-900 MPa) and rubber-like PLCL 50:50 (1.5-12 MPa); a blend has an intermediate stiffness. |
| Gradient | How the hard/soft mix changes inside the cage: axial (soft at the top and bottom faces), radial (soft toward the window), or both. |
| Mixing nozzle / interleave | Two printing methods: one nozzle mixes both materials; or two nozzles print separate hard and soft lines side by side. |
| Distraction d | How far (0.5-1.5 mm) the surgeon spreads the vertebrae before inserting the cage. The cage is made this much taller than the natural gap. |
| Clamp force | The force with which the vertebrae squeeze the cage after the spreader is released. More clamp means more friction grip. |
| Friction coefficient mu | How "grippy" the cage-bone contact is (0.2-0.5). We use the worst case, 0.2, for slip checks. |
| Anti-slip safety factor SF | Available friction (mu x squeezing force) divided by the sideways force trying to slide the cage. SF < 1 means the cage would slide. |
| Tipping | The cage rocks over one edge. Measured by where the net squeezing force acts: at or beyond the cage edge means tipping. |
| Load capacity | The largest compressive force before the first failure sign: cage damage, bone yielding under the cage, too much height loss, or the force peak. |
| Subsidence | The cage sinking into the vertebra because the bone under it yields. |
| ISO 18192-1 | International test standard for wear of permanent artificial discs: the vertebra is forced through set bending/twisting angles with a 100-150 N axial load [14]. |
| Follower load | A compressive force that follows the curve of the spine, mimicking muscle and body weight (here 100 N) [17]. |
| Pure moment | A bending or twisting load without push or pull; 1.5 N m is a standard everyday-level value for the neck [1]. |
| ROM | Range of motion: how many degrees the segment bends under a given load. |
| ICR | Instantaneous centre of rotation: the point the upper vertebra naturally rotates about [15,16]. |
| Mesh / mesh convergence | The element size; convergence = results stop changing when elements get smaller (here < 5 %). |
| Contact penalty | A numerical stiffness that stops two surfaces passing through each other; too low lets them overlap, too high makes the solver fail. |
| Bonded screening interface | A simplification: after insertion the cage is glued to the bone in the model, and the computed glue forces show whether real friction could hold it (Decision: Q47). |
| Element inversion | An element squeezed so much that it turns inside out; the solver must stop. Used to locate the weakest material. |
| Functional failure / lower bound | When the solver stops because a layer is crushed before any formal failure sign, we count it as failure and record the load so far as a minimum value (Decision: Q40). |
| Active learning (AL) | A computer strategy that chooses the next designs to simulate, based on a statistical model (Gaussian process, GP) of all results so far [23,24]. |
| Robust design | A design that performs well over the whole range of uncertain material properties, not just at average values. |
| Preflight | An automatic check before every run that the case was built with the current code and settings (Comm.: PI 2026-10-22). |

# 3. What we are doing now

## 3.1 Goal of the simulations

Question 1 - Does the cage stay in place? It must not slide or tip at any time during the test, because "even a small dislocation is a problem in the neck" (Comm.: PI 2026-10-21)(Decision: Q48, Q51).

Question 2 - Is it strong enough? It must carry at least 150 N, the largest axial force in the motion test (Decision: Q49). A healthy lower neck segment can take about 1164 N [8].

Question 3 - Which design features matter? We compare two cage shapes, three gradients, two printing methods and the distraction range, over the full range of uncertain material values. The aim is a robust design (Decision: Test-matrix strategy 2026-10-20).

## 3.2 How the model is built (conditions and assumptions)

### Spine model

Geometry and tissue properties come from a published, freely available FEBio neck model of a 26-year-old woman [1]. It is idealised and left-right symmetric [6]. Its C3-C4 disc is 5.0 mm high (population average 3.6 +- 1.1 mm) [7].

Bones can deform: a stiff outer shell (cortical bone) around softer spongy bone (trabecular bone), with the original model values [1](Decision: Materials 2026-10-20).

We simulate only the C3-C4 segment, not the whole neck. We checked that this gives the same bending behaviour as the full neck model (within about 1-5 %) (Decision: Single-segment check 2026-10-20).

### Surgery in the model

Removed: the front ligament (ALL), the back ligament (PLL), the nucleus, the front part of the annulus, and the cartilage on the endplates. The bony endplates are kept (Decision: Surgery 2026-10-20).

Kept: facet joints with their capsules, and the ligaments at the back of the spine (Decision: Surgery 2026-10-20).

The leftover side/back annulus ring was first kept. It is now treated as removed, because it made the model 50-100 times too stiff when the vertebrae are spread apart (see Section 3.4) (Decision: PI 2026-10-22 22:05)[21,22].

### Cage

Outline = vertebra outline shrunk by 1 mm. The window is 60 % of the area. The tilt follows the natural endplate angle. Height = natural gap + d, with d from 0.5 to 1.5 mm (Decision: Cage geometry 2026-10-20).

Design A: top and bottom surfaces shaped to match the bone exactly, no pad.

Design B: a simpler flat cage. It is narrower (68 % of A's area) so it fits between the bony ridges on the side of C4 (uncinate processes), and has a 0.6 mm soft pad on each face to fill small gaps (Decision: Q37).

Walls and layers follow what the printer can actually make: 0.4 mm lines and 0.2 mm layers (Decision: Q38).

### Materials (uncertain values are varied, not fixed)

Stiff PLCL 85:15: 250 / 400 / 900 MPa at body temperature [9,28]. Rubber-like PLCL 50:50: 1.5 / 4 / 12 MPa [10,11]. The pad is made of the same family of material (Comm.: Michael).

Blends: stiffness stays low until the hard fraction passes a threshold (20-40 %), then jumps, following measurements on a similar material system [11]. This is an estimate and must be confirmed by our own material samples (Decision: Materials 2026-10-20).

Material model: "Gent" rubber-like law, which stiffens strongly near about 40 % compression. This keeps soft layers from being squashed unrealistically flat in the computer (Decision: Q45)[27]. No test data yet, so the behaviour under very large compression is a stated limitation (Decision: Q43).

Poisson's ratio (how much a material bulges sideways when squeezed): stiff 0.35, soft 0.45. The soft grade is a true elastomer that stretches 4-20 times before breaking [20]; a value of 0.49 was tested and changed results by less than 1 % (Decision: Q42-Q44).

Strength values for judging cage damage: stiff 15 / 25 / 40 MPa, soft 3 / 15 / 30 MPa (Decision: Tests 2026-10-20).

The biodisc is left out of the screening model. It carries less than 1 N, the same for every design, and it caused numerical problems (Decision: PI 2026-10-22 05:20).

### Cage-bone contact

No glue is modelled, only friction (worst case). Friction coefficient mu = 0.2-0.5 (Decision: Interfaces 2026-10-20).

For fast screening, the cage is "glued" to the bone in the model after insertion. We then check whether real friction could supply the glue forces (safety factor SF) (Decision: Q47). Limitation: glue cannot let an edge lift off, so stability is somewhat overestimated. The best designs will be re-checked with true sliding contact (Decision: Q47, Q50).

### Numerical settings

Element size 0.2 mm in the cage and 0.5 mm in the bone surface. Fine enough for fair comparisons. The final designs will be re-run with finer meshes (Decision: Q46, Q53)(Protocol: P8).

Loads are applied in a frame aligned with the disc and centred on the natural rotation centre (ICR) of C3-C4 [15,16,17,18](Decision: PI 2026-10-20 loading frame).

## 3.3 Test procedures

Every simulated case first goes through P2, then one test (P3, P4 or P6). P1 and P8 check the model itself, and P9-P10 organise the many runs.

| No. | Procedure | What happens, in plain words | Why / source |
| --- | --- | --- | --- |
| P1 | Model validation | The whole neck model is bent forward/back, sideways and twisted with 1.5 N m. Its motion and joint forces are compared with the published model and experiments; pass = within 10 %. | [1,2,3,4,5](Decision: Model scope 2026-10-20) |
| P2 | Insertion (every case) | Spread C3 upward by d -> insert the cage (modelled as a slightly shrunk cage that expands back to full height) -> release the spreader so the tissues squeeze the cage. Records clamp force and settling. | (Decision: PI 2026-10-20; PI 2026-10-22) |
| P3 | Compression strength | Push C3 down slowly until the first failure sign: cage damage > 15 %, bone yielding > 10 % in the 3 mm under the cage, height loss > d, force peak, or the soft face layer crushed. | [13](Decision: Tests; Q40; Q41) |
| P4 | ISO 18192-1 motion test (original main test) | The computer forces C3 through set bending (7.5 deg), side-bending (6 deg) and twisting (4 deg) cycles with 100-150 N axial load, at increasing amplitude up to 3x. Designed for permanent artificial discs. | [14,15](Decision: Tests; Q52) |
| P5 | Slip and tipping check | During the test: safety factor SF (slip if < 1), share of time and area that would slip, and tipping (the net force acting at or beyond the cage edge). | (Decision: Q50; Q51) |
| P6 | Everyday-load test (new main test) | 100 N follower load, then 1.5 N m bending forward, backward, left, right and twisting left, right, one at a time. The spine moves as the load dictates (moment control), like daily life during recovery. | [1,17](Decision: PI 2026-10-23 12:00) |
| P7 | Push-out | With 100 N squeezing, push the cage forward or backward until it moves 3 mm. Compared with commercial discs (Mobi-C 142 N, MOVE-C 202 N, Prodisc-C 304 N). | [19](Decision: Tests 2026-10-20) |
| P8 | Mesh and contact checks | Repeat a case with smaller elements and different contact settings; accept if results change < 5 %. | (Decision: Q46; Q53) |
| P9 | Active learning | 48 spread-out starting designs, then the statistical model picks the next ones; stop at 150 designs or when the best result stops improving. | [23,24](Decision: Test-matrix; 2026-10-22 async) |
| P10 | Preflight | Before a run, rebuild the case with the current code and settings and confirm it is identical; otherwise it is blocked. | (Comm.: PI 2026-10-22) |

## 3.4 Results so far

### The model reproduces known data

The published model is reproduced exactly (25 identical solver steps) after one conversion fix (Decision: Validation 2026-10-20)[1].

Whole-neck motion: 18 of 22 values within 10 % of the reference. C3-C4 joint forces within 2 %. The remaining differences lie inside the experimental scatter (Decision: Validation passed, PI 2026-10-20)[1,2,3,4].

The C3-C4 segment alone behaves like the same level inside the whole neck: bending range within 1-5 %, joint forces within 0.3 N (Decision: 2026-10-20).

### The spreading-force error and its fix (most important lesson)

In cadaver tests, spreading C3-C4 by 1.7 mm takes about 44 N [21,22]. Our first model needed 957 N for 1 mm and 4447 N for 1.5 mm, i.e. 50-100 times too stiff (Decision: 2026-10-22 19:55).

We removed one tissue at a time in the model. Without the leftover annulus ring, only 39 N was needed. That ring carried about 96 % of the force. The reference model had been checked only in bending, never in stretching (Decision: 2026-10-22 21:35)[1].

With the ring treated as removed, the spreading force is 9-95 N for d = 0.5-1.5 mm, the same order as cadavers. All earlier strength and clamp numbers were discarded and re-run (Decision: PI 2026-10-22 22:05).

### Screening results: 144 cage variants (status 2026-10-25)

Active learning picked most later variants from design A, because B never met the strength requirement. Each variant was run through compression (P3) and the ISO motion test (P4).

| Result | What we found | What it means |
| --- | --- | --- |
| Strength (P3) | Design A: 108 of 120 >= 150 N, median 261 N. Design B: 0 of 24, median 53 N. | Full, shape-matched contact spreads the load; the narrow B cage with soft pads is too weak. |
| How strength ends | Cage damage first in 76 variants, bone yielding under the cage in 39, solver stop (crushed layer) in 26. | Usually the cage itself, not the bone, limits strength. |
| Clamp after insertion | Median 31 N (7-115 N). | Much lower than daily loads (100-150 N). The tissues squeeze the cage only weakly, so friction grip is small. |
| ISO motion test (P4) | No variant completes it. The best gets through 8 % of the test. Stops: slip 78, crushed soft layer 38, tipping 28. | Under forced motion every design fails early, within the first second of the ramp-up. |
| Why ISO stops | Forced tilting moves the load onto the cage edge; a few soft edge elements are crushed. Longest-lasting: design A, radial gradient, mixing nozzle, stiffer hard phase, higher clamp. | The soft edge layer is the weak spot. The forced-motion test is harsher than the device's real job. |

Because ISO 18192-1 is made for permanent moving discs, the PI decided to replace it by the everyday-load test P6 for screening, and to keep ISO only for the final designs (Decision: PI 2026-10-23 12:00)[14].

### Everyday-load test P6: first results (42 variants)

Design A variants pass the first step (forward bending) in 14 of 21 cases. C3 bends 0.4-3.3 deg at 1.5 N m. No design B variant passes even the initial 100 N load: its soft printed lines are crushed (Decision: 2026-10-23 18:45 physio summary).

Slip warning: in forward bending the safety factor is 0.56-0.91 with the lowest friction (mu = 0.2), i.e. every design A that got this far would slide. SF grows in proportion to mu, so with mu = 0.35 it would be about 1.0-1.6 and with mu = 0.5 about 1.4-2.3. The real cage-bone friction therefore decides the outcome and must be measured.

No run gets through backward bending yet. Near -1.2 N m the solver stalls. This happens for every design and is a numerical problem in the facet joints, not a cage failure. Two alternative facet contact settings removed the crushed elements but still stall (Decision: 2026-10-23 18:45; 21:00).

Test with backward bending limited to 1.0 N m: the run passes backward bending and right side-bending, then stalls in left side-bending (Decision: 2026-10-23 21:00 test).

Decided without a PI reply (auto): soft printed lines crushing under 1.5 N m count as failure; the worst-case friction stays 0.2 (Decision: 2026-10-23 21:00 auto (no reply)).

## 3.5 Next steps

Fix the numerical stall in backward bending so the everyday-load test can run all six directions; then make it the main screening test for active learning (Decision: PI 2026-10-23 12:00)(Protocol: P6).

Ask the PI whether a 1.0 N m backward-bending limit is acceptable meanwhile (Comm.: pending).

Measure the cage-bone friction coefficient (and test blend samples for stiffness and large compression). The slip verdict depends directly on it [11](Decision: Q43).

Add a measure of how well the biodisc is protected (e.g. change of window height, strain in the window), since the biodisc is left out of the screening model (Comm.: PI 2026-10-23 device purpose).

Re-check the best, worst and contrasting designs with true sliding contact (slip in mm, failure at 1 mm) and with push-out in both directions (Decision: Q50, Q51)(Protocol: P7).

Re-run the final designs with finer meshes and on the 59- and 64-year-old models of the same series (Decision: Q46, Q53; Model scope)[1].

Publish a shareable summary web page (Decision: Visualisation).

# 4. Past experience

## 4.1 Earlier pig-spine study (early October 2026)

Set-up: pig C3-C4 geometry; stiff PLCL cages shaped to the bone, one smooth (A) and one with shoe-tread ridges (B); a home-made GPU solver with simple linear materials. Only the two vertebrae were modelled: no ligaments, facet joints or disc remnant [12,18].

Background: in an earlier thesis, printed PLA cages failed at the back because of a bony process [12].

Results: compression strength about 1300 N for both. In repeated compression (ASTM F2077 style) the tread cage resisted slip much longer: smooth slipped 2 mm at 37.5 % of its strength, tread only at 62.5 % [18]. Six-direction loading: little damage, slip <= 0.53 mm.

Lesson: the results looked very good, but they disagreed with the real bench tests, where cages dislocated (Comm.: PI 2026-10-21). Reasons: no soft tissues, no insertion step, an over-stiff material and an unvalidated model. We therefore switched to the validated human model (Comm.: PI 2026-10-20).

## 4.2 What worked well

Validating before adding the cage, so later numbers rest on a checked model (Protocol: P1)[1,2,3,4,5].

A dated decision log: every parameter was discussed with the PI, with a rule for deciding when no reply comes, so simulations never wait (Comm.: PI 2026-10-20).

Comparing with real-world numbers (cadaver spreading force) exposed the biggest model error [21,22].

Changing one thing at a time before adopting any change, e.g. speed settings, Poisson's ratio, the biodisc (0.03 % effect) (Decision: Q44; 2026-10-22).

Computing infrastructure: GPU solver, automatic job scheduling, restart from saved states, automatic pre-run checks and a watchdog [26](Comm.: PI 2026-10-20 "use every free resource").

Finding which elements fail first to tell numerical problems from real material failure (Decision: 2026-10-22 error census).

## 4.3 Failures and what we did about them

| # | Failure | Cause | How we analysed it | What we changed | Did it help? | Outcome |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Facet joints never touched after converting the reference model to the new FEBio version | One contact setting changed meaning between versions (relative vs absolute distance). | Compared step by step with the original author's log. | Converted the setting. | Yes - exact match. | Solved (Decision: Validation 2026-10-20). |
| 2 | Wrong results with the refined bone mesh | A meshing script attached half the disc nodes to loose points not connected to bone. | Checked mesh connectivity. | Fixed the script, marked old results void. | Yes - within 0.3 N of the original. | Solved (Decision: 2026-10-20 VOID). |
| 3 | Software crash (NaN) and stalls | Faults in FEBio for rare numerical cases. | Reproduced and traced in the source code. | Five small source fixes, each verified on known results. | Yes. | Disclosed in the report (Decision: FEBio patches 1-5b). |
| 4 | Results depended on a contact setting (penalty) | Surfaces overlapped about as much as the soft layer really deforms. | Swept the setting over 1000x. | Many contact variants; finally a glued screening model with a friction check. | Partly; the glued model works. | Adopted (Decision: Penalty sweep; Q47). |
| 5 | Soft face layer crushed in every test | The cage face has only 25 % stiff material, below the blend threshold, so it is very soft (about 5.7 MPa). | Tried other Poisson's ratios, element types, material laws. | Rubber-like Gent law; crushing counted as failure. | Partly. | Recognised as a real design weakness (Decision: Q40-Q45). |
| 6 | Computer memory limit, and server stops killed all runs (5 times; the last stop lost about 25 h, 10-24 to 10-25) | Account memory cap; processes do not survive a server stop. | Measured memory per run; read the logs. | Restart from saved states, watchdog, memory-aware scheduling; admin doubled the cap. | Partly - a full server stop still needs a manual restart. | Ongoing risk; check status daily (Decision: 2026-10-21 checkpointing)(Comm.: PI). |
| 7 | Cage overlapped the bone at insertion; many runs failed | The cage was sized on the unspread gap, while the leftover annulus pulled the endplates inward. | Measured endplate movement during spreading. | Insert a shrunk cage that expands; fixed a FEBio fault in that feature. | Yes. | Root cause later found: the annulus ring (Decision: 2026-10-22)(Comm.: PI proposed the shrink-and-expand idea). |
| 8 | Spreading force 50-100x too high (most important) | The leftover annulus ring carried about 96 % of the force; the reference model was never checked in stretching. | Literature comparison; removed one tissue at a time. | Treat the ring as removed. | Yes - cadaver-like forces. | All earlier numbers re-run (Decision: PI 2026-10-22)[21,22]. |
| 9 | Runs built with outdated settings found only after hours | Settings changed while cases were waiting. | - | Automatic pre-run check (preflight). | Yes. | Standard practice (Comm.: PI 2026-10-22). |
| 10 | ISO motion test: every design stops early | Forced tilting crushes the soft cage edge; the test fits permanent discs, not a temporary holder. | Automatic analysis of where and why each run stopped. | Proposed an everyday-load test instead. | Yes - the PI switched the test. | Decided (Decision: PI 2026-10-23 12:00). |
| 11 | Everyday-load test stalls in backward bending | Numerical difficulty in the facet joints near -1.2 N m, the same for every design. | Tried two facet contact variants and a lower moment. | Under investigation. | Not yet. | Open (Decision: 2026-10-23 21:00). |
| 12 | Small human errors (direction labels, a wrong data column, a mis-click) | Manual reading and labelling. | Cross-checked against coordinates and files. | Corrected openly in the decision log. | Yes. | Corrected (Decision: 2026-10-20; Q49; 2026-10-22). |

# 5. Insights

## 5.1 Promising design directions (to be confirmed)

Use the full footprint and match the bone shape (design A). It gives more contact area, keeps the force inside the cage edge (less tipping) and makes the cage stronger. A possible compromise for easier manufacture is a flat cage with a pad of varying thickness made from the patient's CT scan (Decision: Cage geometry future option).

Avoid a very soft layer on the bone-contact faces. Keep the stiff fraction there above the blend threshold, or make the cage soft on the inside and stiff at the edge (radial gradient), which lasted longest (Decision: 2026-10-23 analysis).

Prefer the mixing nozzle to separate hard/soft lines. Separate soft lines are crushed first in every test (Decision: 2026-10-22; 2026-10-23).

Round or reinforce the cage edges, where crushing starts (Decision: Q44).

Add mechanical anchoring, because tissue squeeze and friction are low: tread ridges or teeth (in the pig study ridges raised the slip limit from 37.5 % to 62.5 %), a keel, a front flange, or glue [18](Decision: Interfaces).

Protect against backward movement toward the spinal cord, since the back ligament is removed. In an early model, backward push-out needed only about 14 N, forward over 60 N [19](Decision: 2026-10-20 push-out).

A stiffer hard phase and moderate clamp help; too much clamp crushes the soft layer. The distraction d itself mattered little (Decision: 2026-10-23 analysis).

## 5.2 Proposed simulation standards

Validate the model in every kind of load you will use (bending, stretching, compression), not only one [1,21,22].

Compare key forces with real-world data (cadaver, surgery) before running large batches (Decision: 2026-10-22 19:55).

Check that results do not depend on contact settings and element size, for each result measure separately (Decision: Q46; Q53).

Elements must be no larger than the smallest design feature (0.2 mm layers -> 0.2 mm elements) (Decision: 2026-10-20).

Verify every software fix on known results and disclose it (Decision: Patch verification).

Check every run before it starts; save restart points; watch for stalled or dead runs (Comm.: PI 2026-10-22)(Protocol: P10).

Record stops honestly: as a minimum value plus the state, and say whether the stop is numerical or a real failure (Decision: Q40).

Change one thing at a time and adopt it only if results barely change (Decision: 2026-10-22 speed options).

Choose tests that match the device's job: everyday loads for a temporary holder; harsh standards only as extra checks [14,17](Comm.: PI 2026-10-23).

Keep a dated decision log and mark replaced results clearly (Comm.: PI 2026-10-20).

## 5.3 Other ideas

Add anchoring features (ridge height, direction, spacing; keel) as design variables in active learning (Protocol: P9).

Re-create the earlier bench tests in the model to calibrate it and learn the real dislocation direction (Comm.: PI 2026-10-21).

Run cheaper coarse simulations first and fine ones only where needed (multi-fidelity) [23](Decision: 2026-10-22 pilot).

Test material samples first: blend stiffness, large compression, friction on bone [11](Decision: Q43).

# 6. Take-home messages

Check the model against reality first. One comparison with cadaver data showed the model was 50-100 times too stiff, and two days of results had to be redone [21,22](Decision: PI 2026-10-22).

Full-size, bone-matched cages (A) clearly beat narrow flat cages with soft pads (B): 108/120 vs 0/24 meet the strength requirement.

A very soft surface layer is the common weak point. It fails first in every test.

After surgery the tissues squeeze the cage only weakly, so friction alone may not hold it. The true friction must be measured, and anchoring features are likely needed.

Tests must fit the device's purpose (temporary support and fixation). The everyday-load test now replaces the forced-motion ISO test, once its numerical stall is fixed (Decision: PI 2026-10-23 12:00).

Discipline keeps a large simulation campaign running: pre-run checks, restart points, one change at a time, a decision log.

All numbers so far are comparisons between designs from a simplified screening model. Absolute values need finer meshes, true sliding contact and material tests (Decision: Q46, Q47, Q53, Q43).

# 7. Timeline of key decisions and communications

| Date | Type | Content | Source |
| --- | --- | --- | --- |
| 10-20 | Comm. | Switch from the pig model to a validated human C3-C4 model; every parameter must be discussed, never assumed | PI |
| 10-20 | Decision | Surgery: disc replacement from the front; which tissues are removed and kept; three-step insertion | Surgery |
| 10-20 | Decision | Cage shape (1 mm inset, 60 % window, d 0.5-1.5 mm); designs A and B; B fitted between the uncinate processes | Q37 |
| 10-20 | Decision | Material ranges, blend law, three gradients, two printing methods | Materials |
| 10-20 | Decision | Tests: compression strength, ISO motion test, push-out; active learning with a robust goal | Tests |
| 10-20 | Comm. | Decide yourself if no reply within 30 min; never stop simulations; use every free resource | PI |
| 10-20 | Decision | Model validation passed; single segment is enough | Validation |
| 10-20 | Decision | A solver stop before any failure sign counts as failure with a minimum value; crushed face layer is a failure sign | Q40, Q41 |
| 10-21 | Decision | Material law and settings (Poisson's ratio 0.45, Gent law); limitation stated | Q42-Q45 |
| 10-21 | Decision | Screening meshes; fine meshes for the final designs | Q46, Q53 |
| 10-21 | Decision | Glued screening model with a friction check | Q47 |
| 10-21 | Comm. | Earlier bench tests failed mainly by cage dislocation | PI |
| 10-21 | Decision | Goal: no dislocation first, strength >= 150 N as a requirement; no slip at any time | Q48-Q51 |
| 10-21 | Comm. | Save restart points; admin doubled the memory limit | PI |
| 10-22 | Comm. | Insert a shrunk cage that expands; check every run before it starts | PI |
| 10-22 | Decision | Biodisc left out of the screening model | PI |
| 10-22 | Decision | Leftover annulus ring treated as removed; earlier numbers discarded | PI |
| 10-23 | Comm. | Switch all runs to the corrected (v2) screening model | PI |
| 10-23 | Comm. | BIOCAGE = temporary support and fixation while the biodisc matures | PI |
| 10-23 | Decision | Replace the ISO test by an everyday-load test for screening; ISO kept for final designs | PI 12:00 |
| 10-23 | Comm. | "No idle resources": pack more runs per machine | PI 12:45 |
| 10-23 | Decision | Crushed soft lines under everyday load count as failure; worst-case friction stays 0.2 | auto (no reply) 21:00 |
| 10-25 | Status | Server stop found (runs idle since 10-24 ~11:00); all services restarted | Analyst |

# 8. References

[1] Herron MR, Brockmeyer DL, Dailey AT, Ellis BJ. J Biomech 2020 (FEBio cervical spine finite-element models, 26 / 59 / 64 y.o.; simtk.org/projects/febiocervical).

[2] Panjabi MM, et al. Mechanical properties of the human cervical spine as shown by three-dimensional load-displacement curves. Spine 2001.

[3] Leahy PD, et al. 2007 (cervical range-of-motion data used in the Herron 2020 validation).

[4] Patel VV, et al. 2017 (C3-4 facet joint force, used in the Herron 2020 validation).

[5] Kretzer RM, et al. 2012 (intradiscal pressure, used in the Herron 2020 validation).

[6] Osth J, et al. 2016 (geometry source of the cervical model, 167 cm / 59 kg female).

[7] Lu J, et al. 1999 (population data on cervical disc height).

[8] Moroney SP, et al. 1988 (maximum C4-5 compression about 1164 N).

[9] PLCL 85:15 stiffness at 37 C, wet (379 +- 132 MPa, with filler). PMC9697335.

[10] PLCL 50:50 elastomer stiffness (1.3-12 MPa). PMC8070679.

[11] Duan et al. 2021, stiffness and strength of hard / soft PLCL blends vs blend fraction. PMC8296602.

[12] Lintz M. 2021. PhD thesis, Cornell University (ch. 2 biodisc stiffness 0.03 MPa; ch. 4 posterior failure of PLA cages).

[13] Morgan EF, Keaveny TM. Dependence of yield strain of human trabecular bone on anatomic site. J Biomech 2001.

[14] ISO 18192-1:2011 Implants for surgery - Wear of total intervertebral spinal disc prostheses - Part 1 (parameters cited only).

[15] Anderst W, et al. 2013. Spine 38:E594 (cervical instantaneous centre of rotation).

[16] Amevo B, Worth D, Bogduk N. 1991. Clin Biomech 6:111 (cervical instantaneous axes of rotation).

[17] Patwardhan AG, et al. 2000. Spine (follower load), PMID 10851105.

[18] ASTM F2077 Test Methods for Intervertebral Body Fusion Devices.

[19] ASTM draft F04.25.02.02 (cage push-out); reference values Mobi-C 142 +- 18 N, MOVE-C 202 +- 8 N (100 N preload), Prodisc-C 304 +- 30 N (45 N preload).

[20] PLCL 50:50 elongation at break 388-1974 %. Soft Matter 2020.

[21] Frank E, et al. 1996 (cadaver C3-C4 spreading: 1.7 mm at 44 N, 2.7 mm at 89 N).

[22] Neurosurgery 1996 (12 cadavers, about 2.1-2.4 mm at 89 N; disc tensile failure about 569 N).

[23] Alvi AS, et al. 2019. Asynchronous batch Bayesian optimisation with improved local penalisation (PLAyBOOK). ICML.

[24] Balandat M, et al. 2020. BoTorch: a framework for efficient Monte-Carlo Bayesian optimization. NeurIPS.

[25] Maas SA, Ellis BJ, Ateshian GA, Weiss JA. 2012. FEBio: finite elements for biomechanics. J Biomech Eng.

[26] NVIDIA cuDSS (GPU direct sparse solver).

[27] Gent AN. 1996. A new constitutive relation for rubber. Rubber Chem Technol.

[28] Corbion PURASORB PLC 8516 product data.

# BIOCAGE 頸椎支架：電腦模擬研究

我們在做什麼、學到了什麼、對設計有什麼意義

狀態截至 2026-10-25　｜　人體 C3-C4 有限元素研究　｜　作者：Yi Liu (austonliuyi)　｜　中文版

來源標註說明：藍色上標 [n] 為已發表文獻（清單見第 8 節）；〔決定：…〕為與計畫主持人（PI）確認、並連同日期記錄在專案決定紀錄中的選擇，「Q47」等為紀錄中的問題編號；〔溝通：…〕為 PI 的指示或與 PI 的溝通；〔流程：Pn〕指第 3.3 節的試驗流程 Pn。「auto (no reply)」表示 PI 約 30 分鐘內未回覆，分析者依 PI 訂定的優先順序（1 最接近真實人體生理、2 有文獻或標準支持、3 分析者有根據的建議、4 省時且不太損失準確度）自行決定，事後告知 PI。專有名詞見名詞解釋（第 2.4 節）。

# 1. 一頁摘要

BIOCAGE 是一個小型 3D 列印高分子支架（cage），在切除受損的椎間盤後放在兩節頸椎之間。支架中央的窗口放置實驗室培養的替代椎間盤（biodisc）。支架的任務只是在 biodisc 於體內成熟期間撐住空間、固定不動 〔溝通：PI 2026-10-23〕。先前用實體支架做的試驗顯示，主要問題是支架移位（脫位）〔溝通：PI 2026-10-21〕。因此我們用經過驗證的人體頸椎電腦模型，測試許多支架變化，找出哪些特徵能讓支架保持穩定且夠強。

電腦模型在彎曲方面可信：重現已發表的參考模型與實驗數據，差距約 10 % 以內 [1,2,3,4,5]〔決定：驗證通過，PI 2026-10-20〕。

發現並修正一個重大錯誤：模型一開始把椎骨撐開所需的力，比真實大體頸椎多 50-100 倍；原因是殘留的一圈椎間盤組織，現在依手術實況視為已切除 [21,22]〔決定：PI 2026-10-22〕。

全尺寸、貼合骨面的支架（設計 A）遠優於窄的平面支架加軟墊（設計 B）：設計 A 的 120 個變化中有 108 個能承受至少 150 N；設計 B 的 24 個沒有一個做到。

軟的表面層在每個試驗都是弱點，而且單靠摩擦可能固定不住支架：在日常彎曲負荷、最低合理摩擦係數下，所有撐到該階段的設計都預測會滑動。

試驗要符合裝置的任務：PI 已把原本的運動試驗（為永久人工椎間盤設計）換成日常負荷試驗 〔決定：PI 2026-10-23 12:00〕；但後仰彎曲時的數值問題要先解決，才能用在所有設計上。

# 2. 給新讀者的背景

## 2.1 臨床問題

頸椎有七節椎骨，從頭往下稱為 C1 到 C7。相鄰椎骨之間是椎間盤：外圈是堅韌的纖維環（annulus），中央是膠狀的髓核（nucleus），椎間盤附著在椎骨上下的平面（終板，endplate）。椎間盤後方有兩個小的滑動關節（小面關節，facet joint），脊椎前後各有一條韌帶（前縱韌帶 ALL、後縱韌帶 PLL）。本研究看的是第 3 與第 4 節頸椎之間（C3-C4）。

椎間盤退化壓迫神經時，外科醫師從頸部前方把它切除，然後用硬支架讓兩節椎骨融合，或放入保留活動度的人工椎間盤（ACDR，前路頸椎間盤置換）〔決定：手術 2026-10-20〕。

## 2.2 BIOCAGE 的構想

用 PLCL（一種會慢慢降解的醫用高分子）3D 列印的支架；混合硬的 PLCL 85:15 與像橡膠的 PLCL 50:50，可做出較硬或較軟的區域 [9,10,11,28]。

中央窗口（支架面積的 60 %）放 biodisc，也就是組織工程椎間盤 [12]〔決定：支架幾何 2026-10-20〕。

用途：在 biodisc 於體內成熟期間提供暫時性支撐與固定；不是永久活動的椎間盤，也不是融合器 〔溝通：PI 2026-10-23〕。

主要風險：支架滑動或翻轉移位（脫位）、支架破裂、支架陷入骨頭（下沉）、在 biodisc 成熟前空間塌陷。

## 2.3 為什麼用電腦模擬、怎麼做

有限元素（FE）模擬把骨頭、組織和支架切成許多小塊（元素），解物理方程式，預測力、變形、滑動與損傷。這讓我們能在完全相同的條件下測試數百個支架變化，實體試驗做不到這麼快。模型必須先與真實量測比對（驗證）〔流程：P1〕。每個模擬案例都走同一個順序：建立脊椎節段 → 模擬手術 → 置入支架 → 施加試驗負荷（流程見第 3.3 節）。

我們使用 FEBio 4.13，一個免費的生物力學研究用 FE 程式 [25]。我們修正了它原始碼中五個小錯誤，並確認修正不會改變原本正確的結果〔決定：修補驗證 2026-10-20〕；另外加入顯示卡（GPU）求解器，使每次運算快約 2-4 倍 [26]。

## 2.4 名詞解釋

| 名詞 | 在本報告中的意思 |
| --- | --- |
| C3-C4 節段 / FSU | 第 3、4 節頸椎，連同中間的椎間盤、關節與韌帶（功能性脊椎單元）。 |
| 終板 | 椎骨上下的骨面，支架就放在上面。 |
| 殘留纖維環 | 醫師切除椎間盤後可能留下的外圈纖維組織。 |
| ACDR | 前路頸椎間盤置換：從前方切除椎間盤，放入保留活動度的裝置。 |
| Biodisc | 放在支架窗口中、實驗室培養的組織工程椎間盤；非常軟（0.03 MPa）。 |
| PLCL 硬相 / 軟相 | 硬的 PLCL 85:15（250-900 MPa）與像橡膠的 PLCL 50:50（1.5-12 MPa）；混合物的硬度介於兩者之間。 |
| 梯度 | 支架內硬/軟比例怎麼變化：軸向（上下表面軟）、徑向（靠窗口軟）、或兩者都有。 |
| 混合噴嘴 / 交錯 | 兩種列印方式：一個噴嘴把兩種材料混在一起；或兩個噴嘴並排印出分開的硬線與軟線。 |
| 撐開量 d | 置入支架前醫師把椎骨撐開的距離（0.5-1.5 mm）；支架比原本的間隙高這麼多。 |
| 夾持力 | 撐開器放開後，椎骨夾住支架的力。夾持越大，摩擦抓力越大。 |
| 摩擦係數 μ | 支架與骨頭接觸面的「抓地」程度（0.2-0.5）。滑動判定用最差情況 0.2。 |
| 防滑安全係數 SF | 可用摩擦力（μ × 夾緊力）除以想讓支架滑動的側向力。SF < 1 代表會滑動。 |
| 翻轉 | 支架繞著一邊翻起。以合力作用點判斷：落在支架邊緣或之外即為翻轉。 |
| 承載能力 | 出現第一個失效徵兆前的最大壓力：支架損傷、支架下方骨頭降伏、高度損失過多，或力量達到峰值。 |
| 下沉 | 支架下方骨頭降伏，使支架陷入椎骨。 |
| ISO 18192-1 | 永久人工椎間盤磨耗試驗的國際標準：在 100-150 N 軸力下，強制椎骨做固定角度的彎曲/扭轉 [14]。 |
| Follower load | 沿著脊椎曲線方向的壓力，模擬肌肉與體重（本研究 100 N）[17]。 |
| 純力矩 | 只彎或只扭、不推不拉的負荷；1.5 N·m 是頸椎常用的日常程度數值 [1]。 |
| ROM | 活動範圍：在某負荷下節段彎曲了幾度。 |
| ICR | 瞬時旋轉中心：上位椎骨自然轉動所繞的點 [15,16]。 |
| 網格 / 網格收斂 | 元素大小；收斂 = 元素再變小，結果也不再改變（本研究 < 5 %）。 |
| 接觸罰函數 | 防止兩個表面互相穿過的數值剛度；太低會重疊，太高會讓求解器失敗。 |
| 黏合式篩選介面 | 一種簡化：置入後在模型中把支架「黏」在骨頭上，再用算出的黏合力判斷真實摩擦是否抓得住 〔決定：Q47〕。 |
| 元素反轉 | 元素被壓到內外翻轉，求解器必須停止；可用來找出最弱的材料。 |
| 功能性失效 / 下限值 | 在任何正式失效徵兆出現前，因某層被壓扁而求解器停止，算作失效，並把當時的負荷記為最低值 〔決定：Q40〕。 |
| 主動學習（AL） | 依據目前所有結果建立統計模型（高斯過程 GP），由電腦挑選下一批要模擬的設計 [23,24]。 |
| 穩健設計 | 在整個不確定的材料性質範圍內都表現良好，而不只在平均值表現好。 |
| Preflight（執行前檢查） | 每次執行前自動確認案例是用目前的程式與設定建立的 〔溝通：PI 2026-10-22〕。 |

# 3. 現在在做什麼

## 3.1 模擬的目的

問題 1 —— 支架會不會留在原位？試驗中任何時刻都不可以滑動或翻轉，因為「頸椎即使小幅脫位都是問題」〔溝通：PI 2026-10-21〕〔決定：Q48, Q51〕。

問題 2 —— 夠不夠強？至少要承受 150 N，也就是運動試驗中最大的軸向力 〔決定：Q49〕；健康的下頸椎節段約可承受 1164 N [8]。

問題 3 —— 哪些設計特徵重要？比較兩種支架形狀、三種梯度、兩種列印方式與撐開量範圍，並涵蓋不確定材料值的全部範圍，目標是穩健設計 〔決定：試驗矩陣策略 2026-10-20〕。

## 3.2 模型怎麼建（條件與假設）

### 脊椎模型

幾何與組織性質取自已發表、可免費取得的 FEBio 頸椎模型（26 歲女性）[1]；模型理想化、左右對稱 [6]；C3-C4 椎間盤高 5.0 mm（族群平均 3.6 ± 1.1 mm）[7]。

骨頭可以變形：堅硬的外殼（皮質骨）包著較軟的海綿骨，使用原模型數值 [1]〔決定：材料 2026-10-20〕。

只模擬 C3-C4 一個節段，不是整條頸椎；已確認它的彎曲行為與整條頸椎模型中的同一節段相同（差約 1-5 %）〔決定：單節段檢查 2026-10-20〕。

### 模型中的手術

移除：前縱韌帶（ALL）、後縱韌帶（PLL）、髓核、前方纖維環、終板上的軟骨；保留骨性終板 〔決定：手術 2026-10-20〕。

保留：小面關節與其關節囊、脊椎後方的韌帶 〔決定：手術 2026-10-20〕。

側後方殘留的纖維環原本保留，現在視為已切除，因為它讓模型在撐開時硬了 50-100 倍（見第 3.4 節）〔決定：PI 2026-10-22 22:05〕[21,22]。

### 支架

外形 = 椎骨輪廓往內縮 1 mm；窗口佔面積 60 %；傾角跟隨原本終板角度；高度 = 原本間隙 + d，d 為 0.5-1.5 mm 〔決定：支架幾何 2026-10-20〕。

設計 A：上下表面完全貼合骨頭形狀，沒有墊片。

設計 B：較簡單的平面支架，較窄（A 面積的 68 %），才能放進 C4 兩側的骨脊（鉤突）之間，每面有 0.6 mm 軟墊填補小間隙 〔決定：Q37〕。

壁厚與層厚依照印表機實際做得到的尺寸：0.4 mm 線寬、0.2 mm 層厚 〔決定：Q38〕。

### 材料（不確定的數值會變動，不固定）

硬的 PLCL 85:15：體溫下 250 / 400 / 900 MPa [9,28]；像橡膠的 PLCL 50:50：1.5 / 4 / 12 MPa [10,11]；墊片用同一族材料 〔溝通：Michael〕。

混合物：硬相比例超過門檻（20-40 %）之前硬度都很低，超過後跳升；依據類似材料系統的量測 [11]。這是推估，必須用我們自己的材料樣品確認 〔決定：材料 2026-10-20〕。

材料模型：「Gent」類橡膠定律，在約 40 % 壓縮附近會急劇變硬，避免軟層在電腦中被不真實地壓平 〔決定：Q45〕[27]；目前沒有試驗數據，極大壓縮下的行為列為限制 〔決定：Q43〕。

Poisson 比（材料被壓時往側邊鼓出的程度）：硬相 0.35、軟相 0.45；軟相是真正的彈性體，斷裂前可拉長 4-20 倍 [20]；測試過 0.49，結果改變不到 1 % 〔決定：Q42-Q44〕。

判斷支架損傷用的強度：硬相 15 / 25 / 40 MPa、軟相 3 / 15 / 30 MPa 〔決定：試驗 2026-10-20〕。

篩選模型不含 biodisc：它承受不到 1 N、對每個設計都一樣，且會造成數值問題 〔決定：PI 2026-10-22 05:20〕。

### 支架與骨頭的接觸

不模擬黏著劑，只有摩擦（最差情況）；摩擦係數 μ = 0.2-0.5 〔決定：介面 2026-10-20〕。

為了快速篩選，置入後在模型中把支架「黏」在骨頭上，再檢查真實摩擦是否能提供這些黏合力（安全係數 SF）〔決定：Q47〕。限制：黏合不允許邊緣掀起，所以穩定性會略為高估；最佳設計會再用真正的滑動接觸重新檢查 〔決定：Q47, Q50〕。

### 數值設定

支架元素 0.2 mm、骨表面 0.5 mm，足以做公平比較；最終設計會用更細的網格重跑 〔決定：Q46, Q53〕〔流程：P8〕。

負荷施加在與椎間盤對齊的座標中，以 C3-C4 自然旋轉中心（ICR）為原點 [15,16,17,18]〔決定：PI 2026-10-20 負荷座標〕。

## 3.3 試驗流程

每個模擬案例都先做 P2，再做一個試驗（P3、P4 或 P6）；P1 與 P8 檢查模型本身，P9-P10 負責組織大量運算。

| 編號 | 流程 | 白話說明 | 原因 / 來源 |
| --- | --- | --- | --- |
| P1 | 模型驗證 | 整條頸椎模型以 1.5 N·m 做前後彎、側彎與扭轉，把活動度與關節力和已發表模型及實驗比較；差距 10 % 內為合格。 | [1,2,3,4,5]〔決定：模型範圍 2026-10-20〕 |
| P2 | 置入（每個案例） | 把 C3 往上撐開 d → 置入支架（模擬為稍微縮小、再膨脹回原高度的支架）→ 放開撐開器，讓組織夾住支架；記錄夾持力與沉降。 | 〔決定：PI 2026-10-20；PI 2026-10-22〕 |
| P3 | 壓縮強度 | 慢慢把 C3 往下壓，直到第一個失效徵兆：支架損傷 > 15 %、支架下 3 mm 內骨頭降伏 > 10 %、高度損失 > d、力量峰值、或軟表面層被壓扁。 | [13]〔決定：試驗；Q40；Q41〕 |
| P4 | ISO 18192-1 運動試驗（原本的主試驗） | 電腦在 100-150 N 軸力下，強制 C3 做固定的前後彎（7.5°）、側彎（6°）與扭轉（4°）循環，振幅逐步加大到 3 倍；此標準是為永久人工椎間盤設計的。 | [14,15]〔決定：試驗；Q52〕 |
| P5 | 滑動與翻轉檢查 | 試驗中計算：安全係數 SF（< 1 會滑動）、會滑動的時間與面積比例、翻轉（合力落在支架邊緣或之外）。 | 〔決定：Q50；Q51〕 |
| P6 | 日常負荷試驗（新的主試驗） | 100 N follower load，再依序做前彎、後仰、左右側彎、左右扭轉，各 1.5 N·m；脊椎依負荷自由活動（力矩控制），像恢復期的日常生活。 | [1,17]〔決定：PI 2026-10-23 12:00〕 |
| P7 | 推出試驗 | 在 100 N 夾緊下把支架往前或往後推，直到移動 3 mm；與市售椎間盤比較（Mobi-C 142 N、MOVE-C 202 N、Prodisc-C 304 N）。 | [19]〔決定：試驗 2026-10-20〕 |
| P8 | 網格與接觸檢查 | 用更小的元素與不同接觸設定重跑同一案例；結果變化 < 5 % 才接受。 | 〔決定：Q46；Q53〕 |
| P9 | 主動學習 | 先跑 48 個分散的起始設計，之後由統計模型挑下一批；到 150 個設計或最佳結果不再進步時停止。 | [23,24]〔決定：試驗矩陣；2026-10-22 非同步〕 |
| P10 | Preflight（執行前檢查） | 執行前用目前的程式與設定重建案例並確認完全相同，否則擋下不跑。 | 〔溝通：PI 2026-10-22〕 |

## 3.4 目前的結果

### 模型能重現已知數據

一個版本轉換的設定修正後，已發表模型被完全重現（25 個求解步驟完全相同）〔決定：驗證 2026-10-20〕[1]。

整條頸椎的活動度：22 個數值中 18 個與參考差距在 10 % 內；C3-C4 關節力差距 2 % 內；其餘差距都在實驗數據的分布內 〔決定：驗證通過，PI 2026-10-20〕[1,2,3,4]。

單獨的 C3-C4 節段與整條頸椎中的同一節段行為一致：彎曲範圍差 1-5 %、關節力差 0.3 N 內 〔決定：2026-10-20〕。

### 撐開力錯誤與修正（最重要的教訓）

大體試驗中，把 C3-C4 撐開 1.7 mm 約需 44 N [21,22]；我們最初的模型撐開 1 mm 需 957 N、1.5 mm 需 4447 N，硬了 50-100 倍 〔決定：2026-10-22 19:55〕。

在模型中一次移除一種組織：沒有殘留纖維環時只需 39 N——這圈組織承擔了約 96 % 的力。參考模型只在彎曲下驗證過，從未在拉伸下驗證 〔決定：2026-10-22 21:35〕[1]。

把這圈組織視為已切除後，d = 0.5-1.5 mm 的撐開力為 9-95 N，與大體同數量級。之前所有強度與夾持數字都作廢並重跑 〔決定：PI 2026-10-22 22:05〕。

### 篩選結果：144 個支架變化（2026-10-25 狀態）

因為設計 B 從未達到強度要求，主動學習後來挑的變化大多是設計 A。每個變化都做了壓縮（P3）與 ISO 運動試驗（P4）。

| 結果 | 我們發現 | 代表什麼 |
| --- | --- | --- |
| 強度（P3） | 設計 A：120 個中 108 個 ≥ 150 N，中位數 261 N；設計 B：24 個中 0 個，中位數 53 N。 | 完整、貼合的接觸面能分散負荷；窄的 B 支架加軟墊太弱。 |
| 強度如何結束 | 76 個是支架先損傷、39 個是支架下骨頭降伏、26 個是求解器停止（某層被壓扁）。 | 限制強度的通常是支架本身，不是骨頭。 |
| 置入後夾持力 | 中位數 31 N（7-115 N）。 | 遠低於日常負荷（100-150 N）：組織夾支架的力很小，摩擦抓力也小。 |
| ISO 運動試驗（P4） | 沒有任何變化完成；最好的只完成 8 %。停止原因：滑動 78、軟層壓扁 38、翻轉 28。 | 在強制運動下，所有設計都在起始的第一秒內就失效。 |
| ISO 為何停止 | 強制傾斜把負荷移到支架邊緣，少數軟的邊緣元素被壓扁。撐最久的是：設計 A、徑向梯度、混合噴嘴、較硬的硬相、較大夾持。 | 軟的邊緣層是弱點；強制運動試驗比裝置的真實任務嚴苛。 |

因為 ISO 18192-1 是為永久活動椎間盤設計的，PI 決定篩選改用日常負荷試驗 P6，ISO 只用於最終設計 〔決定：PI 2026-10-23 12:00〕[14]。

### 日常負荷試驗 P6：初步結果（42 個變化）

設計 A 的 21 個變化中有 14 個通過第一步（前彎），1.5 N·m 下 C3 彎 0.4-3.3°；設計 B 連一開始的 100 N 負荷都撐不過，印出的軟線被壓扁 〔決定：2026-10-23 18:45 日常負荷試驗摘要〕。

滑動警訊：前彎時在最低摩擦（μ = 0.2）下安全係數為 0.56-0.91，也就是所有撐到這一步的設計 A 都會滑動。SF 與 μ 成正比：μ = 0.35 時約 1.0-1.6，μ = 0.5 時約 1.4-2.3。所以真實的支架-骨頭摩擦係數決定結果，必須量測。

目前沒有任何案例通過後仰：接近 -1.2 N·m 時求解器停滯。所有設計都一樣，是小面關節的數值問題，不是支架失效；兩種替代的小面接觸設定消除了被壓壞的元素，但仍然停滯 〔決定：2026-10-23 18:45；21:00〕。

把後仰限制在 1.0 N·m 的測試：通過後仰與右側彎，在左側彎停滯 〔決定：2026-10-23 21:00 測試〕。

PI 未回覆而自行決定：1.5 N·m 下印出的軟線被壓扁算作失效；最差摩擦維持 0.2 〔決定：2026-10-23 21:00 auto (no reply)〕。

## 3.5 下一步

解決後仰時的數值停滯，讓日常負荷試驗能跑完六個方向，再把它設為主動學習的主篩選試驗 〔決定：PI 2026-10-23 12:00〕〔流程：P6〕。

詢問 PI：在此之前，後仰暫時限制在 1.0 N·m 是否可接受 〔溝通：待決〕。

量測支架-骨頭摩擦係數（並測試混合材料樣品的硬度與大壓縮行為），滑動判定直接取決於它 [11]〔決定：Q43〕。

加入 biodisc 受保護程度的指標（例如窗口高度變化、窗口內應變），因為篩選模型不含 biodisc 〔溝通：PI 2026-10-23 裝置目的〕。

用真正的滑動接觸重新檢查最好、最差與對照的設計（滑動以 mm 計，1 mm 即失效），並做前後兩方向推出試驗 〔決定：Q50, Q51〕〔流程：P7〕。

最終設計用更細網格重跑，並在同系列的 59 歲與 64 歲模型上重跑 〔決定：Q46, Q53；模型範圍〕[1]。

發佈可分享的摘要網頁 〔決定：視覺化〕。

# 4. 過去的經驗

## 4.1 先前的豬頸椎研究（2026 年 10 月初）

設定：豬 C3-C4 幾何；貼合骨面的硬 PLCL 支架，一個光滑（A）、一個有鞋底狀凸紋（B）；自寫的 GPU 求解器、簡單線性材料；只模擬兩節椎骨，沒有韌帶、小面關節或殘留椎間盤 [12,18]。

背景：先前的博士論文中，列印的 PLA 支架因椎骨骨突而在後方失效 [12]。

結果：兩者壓縮強度都約 1300 N；反覆壓縮（ASTM F2077 方式）中凸紋支架抗滑動久得多：光滑支架在 37.5 % 強度時滑 2 mm，凸紋支架到 62.5 % 才滑 [18]；六方向負荷：損傷很小、滑動 ≤ 0.53 mm。

教訓：結果看起來很好，卻與實體試驗（支架脫位）不符 〔溝通：PI 2026-10-21〕；原因是沒有軟組織、沒有置入過程、材料太硬、模型未經驗證。因此改用經驗證的人體模型 〔溝通：PI 2026-10-20〕。

## 4.2 做得好的地方

加入支架前先驗證，後續數字都建立在檢查過的模型上 〔流程：P1〕[1,2,3,4,5]。

有日期的決定紀錄：每個參數都和 PI 討論，並有 PI 未回覆時的決定規則，模擬不必等待 〔溝通：PI 2026-10-20〕。

與真實世界數字比較（大體撐開力），抓出最大的模型錯誤 [21,22]。

一次只改一件事，確認後才採用，例如速度設定、Poisson 比、biodisc（影響 0.03 %）〔決定：Q44；2026-10-22〕。

運算基礎建設：GPU 求解器、自動排程、從存檔點續跑、執行前自動檢查、看門狗 [26]〔溝通：PI 2026-10-20「利用任何一分資源」〕。

找出最先壞掉的元素，區分數值問題與真實材料失效 〔決定：2026-10-22 錯誤普查〕。

## 4.3 失敗經驗與處理

| # | 失敗經驗 | 原因 | 如何分析 | 做了什麼更改 | 有用嗎 | 結局 |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 參考模型轉成新版 FEBio 後，小面關節從未接觸 | 一個接觸設定在新舊版本意義不同（相對距離 vs 絕對距離）。 | 與原作者的 log 逐步比對。 | 換算該設定。 | 有用——完全一致。 | 已解決 〔決定：驗證 2026-10-20〕。 |
| 2 | 加密骨網格後結果錯誤 | 網格腳本把一半椎間盤節點接到沒連上骨頭的散點。 | 檢查網格連接性。 | 修正腳本，舊結果宣告作廢。 | 有用——與原模型差 0.3 N 內。 | 已解決 〔決定：2026-10-20 作廢〕。 |
| 3 | 程式當機（NaN）與停滯 | FEBio 在少見數值情況下的錯誤。 | 重現並追到原始碼。 | 五個小修補，各自在已知結果上驗證。 | 有用。 | 在報告中揭露 〔決定：FEBio 修補 1-5b〕。 |
| 4 | 結果受接觸設定（罰函數）影響 | 表面重疊量和軟層真實變形差不多大。 | 把設定掃過 1000 倍範圍。 | 試過多種接觸方式，最後採用黏合篩選模型加摩擦檢查。 | 部分有用；黏合模型可行。 | 已採用 〔決定：罰函數掃描；Q47〕。 |
| 5 | 每個試驗中軟表面層都被壓扁 | 支架表面只有 25 % 硬相，低於混合門檻，所以非常軟（約 5.7 MPa）。 | 試過其他 Poisson 比、元素類型、材料定律。 | Gent 類橡膠定律；壓扁算作失效。 | 部分有用。 | 判定為真實的設計弱點 〔決定：Q40-Q45〕。 |
| 6 | 電腦記憶體上限，以及伺服器停機殺掉所有運算（5 次；最近一次 10-24 到 10-25 損失約 25 小時） | 帳號記憶體上限；程序撐不過伺服器停機。 | 量測每個運算的記憶體；讀 log。 | 從存檔點續跑、看門狗、記憶體感知排程；管理員把上限加倍。 | 部分有用——整台伺服器停機仍需手動重啟。 | 持續風險；每天檢查狀態 〔決定：2026-10-21 checkpoint〕〔溝通：PI〕。 |
| 7 | 置入時支架與骨頭重疊，大量運算失敗 | 支架依未撐開的間隙設計，而殘留纖維環把終板往內拉。 | 量測撐開時終板的移動。 | 置入縮小後再膨脹的支架；修正 FEBio 這個功能的錯誤。 | 有用。 | 後來找到根本原因：纖維環 〔決定：2026-10-22〕〔溝通：PI 提出縮小再膨脹的構想〕。 |
| 8 | 撐開力高了 50-100 倍（最重要） | 殘留纖維環承擔約 96 % 的力；參考模型從未在拉伸下驗證。 | 文獻比對；一次移除一種組織。 | 把這圈組織視為已切除。 | 有用——力量與大體相近。 | 之前所有數字重跑 〔決定：PI 2026-10-22〕[21,22]。 |
| 9 | 跑了幾小時才發現案例用的是舊設定 | 案例等待期間設定改變了。 | — | 執行前自動檢查（preflight）。 | 有用。 | 成為標準作法 〔溝通：PI 2026-10-22〕。 |
| 10 | ISO 運動試驗：每個設計都提早停止 | 強制傾斜壓壞軟的支架邊緣；這個試驗適合永久椎間盤，不適合暫時性支架。 | 自動分析每個運算在哪裡、為何停止。 | 提出改用日常負荷試驗。 | 有用——PI 已換試驗。 | 已決定 〔決定：PI 2026-10-23 12:00〕。 |
| 11 | 日常負荷試驗在後仰時停滯 | 在約 -1.2 N·m 時小面關節的數值困難，每個設計都一樣。 | 試過兩種小面接觸設定與較低力矩。 | 調查中。 | 尚未。 | 未解決 〔決定：2026-10-23 21:00〕。 |
| 12 | 小的人為錯誤（方向標籤、讀錯資料欄、誤點） | 人工判讀與標註。 | 對照座標與檔案交叉檢查。 | 在決定紀錄中公開更正。 | 有用。 | 已更正 〔決定：2026-10-20；Q49；2026-10-22〕。 |

# 5. 啟發

## 5.1 可能優良的設計方向（尚待確認）

用完整足跡並貼合骨頭形狀（設計 A）：接觸面積更大，合力留在支架邊緣以內（較不會翻轉），支架也更強。若要較易製造，可折衷為平面支架加上依病人 CT 製作、厚度不一的墊片 〔決定：支架幾何 未來選項〕。

接觸骨頭的表面避免非常軟的層：該處硬相比例保持在混合門檻以上，或做成內軟外硬（徑向梯度）——它撐得最久 〔決定：2026-10-23 分析〕。

優先用混合噴嘴，而非分開的硬/軟線：分開的軟線在每個試驗都最先被壓扁 〔決定：2026-10-22；2026-10-23〕。

把支架邊緣倒圓或加強，壓扁從邊緣開始 〔決定：Q44〕。

加入機械固定，因為組織夾持與摩擦都很小：凸紋或齒（豬研究中凸紋把滑動門檻從 37.5 % 提高到 62.5 %）、龍骨、前方翼板或黏著劑 [18]〔決定：介面〕。

防止往後（往脊髓方向）移動，因為後縱韌帶已切除；早期模型中往後推出只需約 14 N，往前超過 60 N [19]〔決定：2026-10-20 推出〕。

較硬的硬相與適度夾持有幫助；夾持太大會壓扁軟層；撐開量 d 本身影響不大 〔決定：2026-10-23 分析〕。

## 5.2 建議的模擬標準

模型要在每一種會用到的負荷下驗證（彎曲、拉伸、壓縮），不能只驗一種 [1,21,22]。

在跑大量案例之前，把關鍵力量與真實數據（大體、手術）比較 〔決定：2026-10-22 19:55〕。

確認結果不受接觸設定與元素大小影響，而且每個結果指標要分開檢查 〔決定：Q46；Q53〕。

元素不能大於最小的設計特徵（0.2 mm 層 → 0.2 mm 元素）〔決定：2026-10-20〕。

每個軟體修補都要在已知結果上驗證並揭露 〔決定：修補驗證〕。

每次執行前檢查；保存續跑點；監看停滯或死掉的運算 〔溝通：PI 2026-10-22〕〔流程：P10〕。

誠實記錄停止：記為最低值加上當時狀態，並說明是數值問題還是真實失效 〔決定：Q40〕。

一次只改一件事，結果幾乎不變才採用 〔決定：2026-10-22 速度選項〕。

選擇符合裝置任務的試驗：暫時性支架用日常負荷；嚴苛標準只當額外檢查 [14,17]〔溝通：PI 2026-10-23〕。

保存有日期的決定紀錄，清楚標示被取代的結果 〔溝通：PI 2026-10-20〕。

## 5.3 其他想法

把固定特徵（凸紋高度、方向、間距；龍骨）加入主動學習的設計變數 〔流程：P9〕。

在模型中重現先前的實體試驗，用來校準模型並找出真實的脫位方向 〔溝通：PI 2026-10-21〕。

先跑便宜的粗略模擬，只在需要的地方跑細緻模擬（多保真度）[23]〔決定：2026-10-22 試行〕。

先做材料樣品試驗：混合硬度、大壓縮、與骨頭的摩擦 [11]〔決定：Q43〕。

# 6. Take-home messages

先拿模型和真實比對。一次與大體數據的比較就發現模型硬了 50-100 倍，兩天的結果必須重做 [21,22]〔決定：PI 2026-10-22〕。

全尺寸、貼合骨面的支架（A）明顯勝過窄的平面支架加軟墊（B）：達到強度要求的比例為 108/120 對 0/24。

非常軟的表面層是共同弱點，在每個試驗都最先失效。

手術後組織夾支架的力很小，單靠摩擦可能固定不住；必須量測真實摩擦係數，很可能需要機械固定特徵。

試驗要符合裝置用途（暫時支撐與固定）：日常負荷試驗在解決數值停滯後，將取代強制運動的 ISO 試驗 〔決定：PI 2026-10-23 12:00〕。

紀律讓大規模模擬持續運作：執行前檢查、續跑點、一次改一件事、決定紀錄。

目前所有數字都是簡化篩選模型中設計之間的比較；絕對值需要更細網格、真正滑動接觸與材料試驗確認 〔決定：Q46, Q47, Q53, Q43〕。

# 7. 重要決定與溝通時間軸

| 日期 | 類型 | 內容 | 出處 |
| --- | --- | --- | --- |
| 10-20 | 溝通 | 從豬模型改為經驗證的人體 C3-C4 模型；每個參數都要討論，不可自行假設 | PI |
| 10-20 | 決定 | 手術：前路椎間盤置換；移除與保留的組織；三步驟置入 | 手術 |
| 10-20 | 決定 | 支架形狀（內縮 1 mm、窗口 60 %、d 0.5-1.5 mm）；設計 A 與 B；B 放在鉤突之間 | Q37 |
| 10-20 | 決定 | 材料範圍、混合定律、三種梯度、兩種列印方式 | 材料 |
| 10-20 | 決定 | 試驗：壓縮強度、ISO 運動試驗、推出試驗；以穩健為目標的主動學習 | 試驗 |
| 10-20 | 溝通 | 30 分鐘未回覆就自行決定；模擬不停；利用所有空閒資源 | PI |
| 10-20 | 決定 | 模型驗證通過；單一節段即可 | 驗證 |
| 10-20 | 決定 | 在任何失效徵兆前求解器停止算作失效並記最低值；表面層壓扁為失效徵兆 | Q40, Q41 |
| 10-21 | 決定 | 材料定律與設定（Poisson 比 0.45、Gent 定律）；列出限制 | Q42-Q45 |
| 10-21 | 決定 | 篩選用網格；最終設計用細網格 | Q46, Q53 |
| 10-21 | 決定 | 黏合篩選模型加摩擦檢查 | Q47 |
| 10-21 | 溝通 | 先前實體試驗主要失敗為支架脫位 | PI |
| 10-21 | 決定 | 目標：防脫位優先，強度 ≥ 150 N 為要求；任何時刻都不可滑動 | Q48-Q51 |
| 10-21 | 溝通 | 保存續跑點；管理員把記憶體上限加倍 | PI |
| 10-22 | 溝通 | 置入縮小再膨脹的支架；每次執行前先檢查 | PI |
| 10-22 | 決定 | 篩選模型不含 biodisc | PI |
| 10-22 | 決定 | 殘留纖維環視為已切除；之前數字作廢 | PI |
| 10-23 | 溝通 | 所有運算改用修正後（v2）的篩選模型 | PI |
| 10-23 | 溝通 | BIOCAGE = biodisc 成熟期間的暫時支撐與固定 | PI |
| 10-23 | 決定 | 篩選改用日常負荷試驗取代 ISO；ISO 保留給最終設計 | PI 12:00 |
| 10-23 | 溝通 | 「不要有閒置資源」：每台機器塞更多運算 | PI 12:45 |
| 10-23 | 決定 | 日常負荷下軟線壓扁算作失效；最差摩擦維持 0.2 | auto (no reply) 21:00 |
| 10-25 | 狀態 | 發現伺服器停機（10-24 約 11:00 起運算閒置）；所有服務已重啟 | 分析者 |

# 8. 參考文獻

[1] Herron MR, Brockmeyer DL, Dailey AT, Ellis BJ. J Biomech 2020（FEBio 頸椎有限元素模型，26 / 59 / 64 歲；simtk.org/projects/febiocervical）。

[2] Panjabi MM, et al. Mechanical properties of the human cervical spine as shown by three-dimensional load-displacement curves. Spine 2001.

[3] Leahy PD, et al. 2007（頸椎活動度數據，Herron 2020 驗證所用）。

[4] Patel VV, et al. 2017（C3-4 小面關節力，Herron 2020 驗證所用）。

[5] Kretzer RM, et al. 2012（椎間盤內壓，Herron 2020 驗證所用）。

[6] Östh J, et al. 2016（頸椎模型幾何來源，167 cm / 59 kg 女性）。

[7] Lu J, et al. 1999（頸椎椎間盤高度族群數據）。

[8] Moroney SP, et al. 1988（C4-5 最大壓縮約 1164 N）。

[9] PLCL 85:15 37 °C 濕態硬度（379 ± 132 MPa，含填料）。PMC9697335.

[10] PLCL 50:50 彈性體硬度（1.3-12 MPa）。PMC8070679.

[11] Duan 等 2021，硬/軟 PLCL 混合物硬度與強度隨比例的變化。PMC8296602.

[12] Lintz M. 2021. Cornell University 博士論文（第 2 章 biodisc 硬度 0.03 MPa；第 4 章 PLA 支架後方失效）。

[13] Morgan EF, Keaveny TM. Dependence of yield strain of human trabecular bone on anatomic site. J Biomech 2001.

[14] ISO 18192-1:2011 Implants for surgery — Wear of total intervertebral spinal disc prostheses — Part 1（只引用參數）。

[15] Anderst W, et al. 2013. Spine 38:E594（頸椎瞬時旋轉中心）。

[16] Amevo B, Worth D, Bogduk N. 1991. Clin Biomech 6:111（頸椎瞬時旋轉軸）。

[17] Patwardhan AG, et al. 2000. Spine（follower load），PMID 10851105.

[18] ASTM F2077 Test Methods for Intervertebral Body Fusion Devices.

[19] ASTM draft F04.25.02.02（支架推出試驗）；參考值 Mobi-C 142 ± 18 N、MOVE-C 202 ± 8 N（100 N 預壓）、Prodisc-C 304 ± 30 N（45 N 預壓）。

[20] PLCL 50:50 斷裂伸長率 388-1974 %。Soft Matter 2020.

[21] Frank E, et al. 1996（大體 C3-C4 撐開：1.7 mm @ 44 N、2.7 mm @ 89 N）。

[22] Neurosurgery 1996（12 具大體，約 2.1-2.4 mm @ 89 N；椎間盤拉伸破壞約 569 N）。

[23] Alvi AS, et al. 2019. Asynchronous batch Bayesian optimisation with improved local penalisation (PLAyBOOK). ICML.

[24] Balandat M, et al. 2020. BoTorch: a framework for efficient Monte-Carlo Bayesian optimization. NeurIPS.

[25] Maas SA, Ellis BJ, Ateshian GA, Weiss JA. 2012. FEBio: finite elements for biomechanics. J Biomech Eng.

[26] NVIDIA cuDSS（GPU 直接稀疏求解器）。

[27] Gent AN. 1996. A new constitutive relation for rubber. Rubber Chem Technol.

[28] Corbion PURASORB PLC 8516 產品資料。